## Slide 1
Gripping Force Comparative Analysis

## Slide 2
Reference Paper Analysis
  ▶ Mucksavage, Phillip, et al. "Differences in grip forces among various robotic instruments and da Vinci surgical platforms." Journal of endourology 25.3 (2011): 523-528.
▶ Measurement of gripping force according to da Vinci system and its end-tool.
▶ Methond of Measurement
  Install Load Cell and build housing to bite (hold) end-tool.
  Using the Bulldog Clamp, compare the measurements of ①directly gripping the Load Cell and ② Gripping the Load Cell after installing it in the housing to derive a conversion formula.
  After gripping the mount with the end-tool, measure the load acting on the Load Cell, and derive the actual gripping force through the conversion formula.

## Slide 3
▶ Main measurement results
- Gripping Force of da Vinci’s Needle Driver : About 15~22N (newton)
Reference Paper Analysis
Maximum 22N

## Slide 4
Reference Paper Analysis
▶ Measurement device dimension estimation and force covernsion
  Estimation of the main dimensions on the photo from actual measurements of Da Vinci end-tools.
3.6mm
(Actual Measurement)
Induced Jaw Length 20.5mm
(Da Vinci data: 20mm, About 2.5% error
Jaw Contact Point: Jaw mounting position 3mm
 Jaw contact point Pivot Length : (Jaw Length) -3mm
  Calculation of Gripping Force Torque on Da Vinci Needle Driver
 | Large Needle Driver | Large SutureCut Needle Driver
Measuring Force (Max) [N] | 20 | 22
Contact point Force [N] | 7.8 | 8.6
Jaw Length [mm] | 10 | 11
Jaw contact point pivot Length [mm] | 7 | 8
Jaw Torque [N mm] | 54.6 | 38.8
Load Cell mounting position
Contact Point Pivot Length26.7mm
Load Cell Pivot Length 10.4mm

## Slide 5
Comparative Analysis with LIVSMED’s Products
▶ Comparison with quality standard
  Main dimensions of LIVSMED NH01, NH02, and Gripping Force measuring equipment
Length of load cell pivot 30mm
Length of contact point pivot 60mm
Jaw contact point pivot
NH01: 6.5mm
NH02: 6.0mm

## Slide 6
Comparative Analysis with LIVSMED’s Products
▶ Comparison with quality standard
  Quality standard Gripping Force of NH01, NH02: Based on Load Cell Measurements 105N ~ 110N
  Calculation of Gripping Force Torque and comparison with Da Vinci End-tool
 | NH01 | NH02
Quality Standard Force (Max.) [N] | 110 | 110
Contact point Force [N] | 55 | 55
Jaw contact point pivot Length [mm] | 6.5 | 6.0
Jaw Torque [N mm] | 357.5 | 330
 | Da Vinci |  | LivsMed | 
 | Large Needle Driver | Large SutureCut Needle Driver | NH01 | NH02
Jaw Torque [N mm] | 54.6 | 38.8 | 357.5 | 330
 LivsMed NH01, NH02 shows approximately 7 to 8 times stronger Gripping Force Torque than Da Vinci Need Driver.
However, in the paper that tested with Da Vinci, due to the structure of the sensor housing, a value lower than the actual Gripping Force was measured by the sensor.
Therefore, the difference in Gripping Force Torque is not up to 7 to 8 times.